VGC/VOC

Precision Counting Scale

Operation Manual

Hao Han Industry Company

February 2015

The information contained here is the property of Hao Han Industry Company and may not be reproduced or used without prior authorization by contract or written permission. Please check with your dealer to ensure the operation manual is the latest version.

TABLE OF CONTENT

- BEFORE USE FOR THE FIRST TIME

- INTRODUCTION

- PANEL AND KEYBOARD

- INTERNAL FUNCTIONS

- INSTRUCTION TO USE

- DAILY CARE AND MAINTENANCE

APPENDIX:

ERROR CODES

DISPLAY SEGMENT DEFINATION

- BEFORE USE FOR THE FIRST TIME

- Thank you for purchasing a HHi product, please read this instruction carefully before start using. You shall find the following items inside the box:

VGC/VOC scale x 1

Switch power adaptor x 1

Operation manual x 1

If you don’t find all the items as described above, please contact your dealer for further assistance.

- Place the scale in a stable environment and adjust the adjustable feet to a level position by referring to the bubble level

- The scale is equipped with rechargeable battery, please charge it until the charging indicator turns green to ensure the best battery performance

- Always use factory supplied power adaptor for charging purpose, an unapproved power adaptor may shorten the battery life and further damages the electronic components

- Warm up the scale for 15 minutes before use for the first time

- Make sure the capacity and division setup on the display when switching on is the same as the data plate

- Avoid to use the scale when the environment is with strong wind, vibration and strong magnetic field

- Use it within the recommend temperature range and avoid wet environment

- Recharge the battery immediately when battery status indicator starts flashing

- Use damp cloth to clean the scale if necessary, excessive moisture is prohibited

- When not using, switch off the power and keep the scale in a dry and cool place

- INTRODUCTION

  - Specification

3000 divisions

| Model No. | Capacity (Max) | Readability (e) |
| --- | --- | --- |
| 1500g | 1500g | 0.5g |
| 3000g | 3000g | 1g |
| 6000g | 6000g | 2g |
| 15kg | 15kg | 5g |
| 30kg | 30kg | 10g |

3000 divisions dual range

| Model No. | Capacity (Max) W1/W2 | Readability (e) e1/e2 |
| --- | --- | --- |
| 3000g | 1500g/3000g | 0.5g/1g |
| 6000g | 3000g/6000g | 1g/2g |
| 15kg | 6kg/15kg | 2g/5g |
| 30kg | 15kg/30kg | 5g/10g |

6000 divisions

| Model No. | Capacity (Max) | Readability (e) |
| --- | --- | --- |
| 3000g | 3000g | 0.5g |
| 6000g | 6000g | 1g |
| 12kg | 12kg | 2g |
| 30kg | 30kg | 5g |

| LCD Digits | VGC: Weight, Unit Weight– 5 digits, Total Count – 6 digits VOC: Weight, Unit Weight, Total Count – 6 digits |
| --- | --- |
| Power-on Zero Range | ±10% of Max |
| Zero Range | ±2% of Max |
| Tare Range | Full Range Tare |
| Platter Size | 230 x 245mm with Stainless Steel Cover |
| Power Source | External AC Adaptor: AC100~240V, DC9V/0.6A Built-in Rechargeable Battery: 6V/4Ah |
| Operating Temperature | -10oC / 40 oC (14 oF / 104 oF) Non-condensed. R.H.≦ 85% |

- PANEL AND KEYBOARD

3.1 Panel

- ZERO INDICATOR

It appears when scale is at zero weight status

- NEGATIVE INDICATOR

It appears when weighing result is below true zero

- NET INDICATOR

It appears when tare function is engaged and weight shown is the net weight

- PRESET TARE INDICATOR

It appears when a preset tare is engaged

- W1 INDICATOR

This indicates the scale is set with dual range, and the weight is within the first range Max1/e1

- W2 INDICATOR

This indicates the scale is set with dual range, and the weight is within the second range Max2/e2

- M+ INDICATOR

This indicates memory contains data

- TOTAL INDICATOR

This indicator appears when value shown on the display is the total saved value

- BATTERY STATUS

This indicates the power level of the rechargeable battery, and should be charged immediately when the sign starts flashing to prolong the time of rechargeable battery

- LOW UNIT WEIGHT INDICATOR

This indicator appears when the fix unit price function is engaged, meaning the unit price enter can only be cancelled manually

- LOW SAMPLE SIZE INDICATOR

This indicator appears when pay & change function is engaged and as a reminder for user to enter the amount receives from customer

- CHARGING STATUS

Red light indicates AC power is supplied and charging is in process.

Green light indicates charging completed

3.2 Keyboard

| ON/OFF | Press this key to turn on the scale, or press and hold this key to turn off the scale |
| --- | --- |
| zero | Press this key to manually set the display value to zero, please refer to specification for manual zero range |
| TARE | 1. Press this key to tare the weight off from the platter 2. Enter numeric value and press this key to preset tare |
| CL | Press this key to clear the unit price entered wrongly, or to escape and returning to weighing mode |
| M+ | Press this key to save the weighing result to memory |
| MR | Press this key to recall the total accumulated weighing results from memory |
| WT. SET | Press this key to manually enter the unit piece weight |
| Q’TY SET | Press this key to enter the number of piece of the load |
| CHK | Press this key to set upper limit for total quantity count or weight |
| 0 ~ 9 | Use these numerical keys to key in price and quantity |

- INTERNAL FUNCTION

Below table defines all the internal function, and kindly note some functions are for dealer only. DO NOT try to open the scale or break with the hardware seal when the unit is metrological approved. Please contact your dealer for more information.

| DISPLAY | FUNCTION SETTING | ACCESS CONDITION |
| --- | --- | --- |
| F 0 | Zero tracking range | Dealer only when LOCKED |
| F 1 | 1. Analog to digital offset value 2. Function key test | Open |
| F 2 | 1. LCD segment check and software version information 2. Adjusting backlight brightness | Open |
| F 3 | Scale configuration setup | Dealer only when LOCKED |
| F 4 | Auto power off | Dealer only when LOCKED |
| F 5 | Transmission protocol for RS-232/RS-485 | Open |
| F 6 | 1. Digital motion filter 2. Animal weighing | Open |
| F 7 | transmission mode & data format for RS-232/RS-485 | Open |
| F 9 | Auto tare | Open |
| F 16 | Turning on/off real time clock | Open |
| F 17 | Time and date setup | Open |
| F 18 | Buzzer sound | Open |
| F 20 | Manufacturing site gravity value setup | Dealer only when LOCKED |
| F 21 | Final destination gravity value setup | Dealer only when LOCKED |
| F 38 | Backlight timing | Open |
| F 99 | High resolution verification mode | Dealer only when LOCKED |
| F 100 | Reset to factory default | Dealer only when LOCKED |

  - Function DEFINITION AND Setup

4.1.1 LCD Segment Check and Backlight Adjustment

Under this function all segments and indicators will be in flashing mode as well as backlight for user to check if function properly. Backlight brightness also has 5 level selections for optimal viewing.

- Scale is off

- Press and hold TARE and press ON/OFF , display shows F 1

- Continue pressing TARE until display shows F2

- Press ZERO to enter, all LCD segments including backlight starts flashing. The display will change between full segment check and software version with buzzer sound for functionality check. Display will also show the software version under this function.

- Press MR to adjust the backlight brightness. There are five settings, press ZERO until the desire brightness is achieved

- When ZERO is pressed, display shows F3

- Press ZERO to enter the next function, or press M+ to exit. When M+ is pressed, display will show SAVE, user can either press ZERO for yes to save the changes made or press CL for not to save the changes.

CAUTION: BEFORE LEAVING THE INTERNAL SETTING, ALL CHANGES MUST BE SAVED OTHERWISE THE SCALE WILL STILL FUNCTION UNDER THE OLD SETTING.

4.1.2 Offset Value & Function Key Test

Under this function scale will show the offset value and be able to check the function of each key.

- Scale is off

- Press and hold TARE and press ON/OFF , display shows F 1

- Press ZERO to enter, display will show the offset value in the TOTAL window. Press any digit key and display will show KEYid and the name of the key you are pressing.

- Press ZERO to exit the current function, display shows F2

- Press ZERO to enter the next function, or press M+ to exit.

4.1.3 Digital Motion Filtering

Under this function, scale will compensate the environmental vibration or unstable weighing object and deliver the correct weighing result.

- Scale is off

- Press and hold TARE and press ON/OFF, display shows F1

- Continue pressing TARE until display shows F6

- Press ZERO to enter, and display shows FiLt0

- Press MR to select the desire setting, there are two levels in the setting

- Press ZERO to confirm and exit the current function, display shows F7

- Press ZERO to enter the next function, or press M+ to exit. When M+ is pressed, display will show SAVE, user can either press ZERO for yes to save the changes made or press CL for not to save the changes.

4.1.4 Auto Power Off

Under this function, user can select auto power off timing among 5 minutes, 10 minutes, 15 minutes and 20 minutes or disable the function.

- Scale is off

- Press and hold TARE and press ON/OFF, display shows F1

- Continue pressing TARE until display shows F4

- Press ZERO to enter, and display shows 5.oFF

- Press MR to select the desire setting. There are 5 settings, the numerical digit means the minutes to power off when scale idles. ALL.on means the Scale will never automatically turn off without manual intervention.

- Press ZERO to confirm and exit the current function, display shows F5

- Press ZERO to enter the next function, or press M+ to exit. When M+ is pressed, display will show SAVE, user can either press ZERO for yes to save the changes made or press CL for not to save the changes.

4.1.5 Backlight Timing

Under this function, user can select when to turn off backlight.

- Scale is off

- Press and hold TARE and press ON/OFF, display shows F1

- Continue pressing TARE until display shows F38

- Press ZERO to enter, and display shows bL.10

- Press MR to select the desire setting. There are 3 settings, bL.on for always on, bL.oFF for always off and bL.10 means the backlight will automatically turn off after 10 seconds when scale idles.

- Press ZERO to confirm and exit the current function, display shows F99

- Press ZERO to enter the next function, or press M+ to exit. When M+ is pressed, display will show SAVE, user can either press ZERO for yes to save the changes made or press CL for not to save the changes.

4.1.6 Buzzer Sound

Under this function, user can select buzzer sound on or off.

- Scale is off

- Press and hold TARE and press ON/OFF, display shows F1

- Continue pressing TARE until display shows F18

- Press ZERO to enter, and display shows Sd.on

- Press MR to select between Sd.on or Sd.oFF

- Press ZERO to confirm and exit the current function, display shows F19(or Fx)

- Press ZERO to enter the next function, or press M+ to exit. When M+ is pressed, display will show SAVE, user can either press ZERO for yes to save the changes made or press CL for not to save the changes.

4.1.7 Auto Tare

Under this function, user can enable or disable auto tare function of the first weight or continuous taring the weight.

- Indicator is off

- Press and hold TARE and press ON/OFF, display shows F1

- Continue pressing TARE until display shows F9

- Press ZERO to enter, and display shows tr.oFF

- Press MR to select between tr.on or tr.oFF

- Press ZERO to confirm and exit the current function, display shows F16(or Fx)

- Press ZERO to enter the next function, or press M+ to exit. When M+ is pressed, display will show SAVE, user can either press ZERO for yes to save the changes made or press CL for not to save the changes.

4.1.8 Transmission Protocol

Under this function, user can select the transmission interface, speed, mode and format for RS-232 and also the communication protocol.

- Scale is off

- Press and hold TARE and press ON/OFF, display shows F1

- Continue pressing TARE until display shows F5

| When display shows F5, press ZERO to enter the function |  |
| --- | --- |
| First level selection: | nonE / rS232 |
| Second level selection: | bAUd / ProtL |
| Selection under bAud: | 4800 / 9600 / 19200 / 38400 / 115200bps |
| Selection under ProtL: | n.8.1 / E.7.1 / o.7.1 / E.8.1 |
| use MR key to select, M+ to change, ZERO key to confirm |  |

- Press ZERO to confirm and exit the current function, display shows F6

- Press ZERO to enter the next function, or press M+ to exit. When M+ is pressed, display will show SAVE, user can either press ZERO for yes to save the changes made or press CL for not to save the changes.

Terminology:

nonE: no data output 				rS232: data output through RS-232 only

bAud: baud rate 					ProtL: transmission protocol

trAns: transmission type				ForMt: transmission format

4.1.8.1 Transmission Mode

Under this function, user can select the transmission interface, speed, mode and format for RS-232 or RS485, and also the communication protocol.

- Indicator is off

- Press and hold TARE and press ON/OFF, display shows  F1

- Continue pressing TARE until display shows  F7

| Selection under trAnS: | StrM / Manu / CMd / Auto |
| --- | --- |
| Selection under ForMt: | SEr.1 / SEr.2 |
| use ZERO key to select, MODE key to confirm |  |

- Press MODE to confirm and exit the current function, display shows  F8

- Press MODE to enter the next function, or press ZERO to exit. When ZERO is pressed, display will show SAVE, user can either press MODE for yes to save the changes made or press ZERO for not to save the changes.

Terminology:

trAnS: transmission mode     ForMt: result format

Sample of print data:

StrM: continue sending data when weighing result stables (only support Ser.1 format)

10.05kg

-10.05kg

6.00kg

0.00kg

……………..

MAnU: data is sent once when manually press M+, and send total value when pressing MC

When format setting is Ser.1, data is sent once when pressing M+, but does not support print total

10.05kg

-10.05kg

6.00kg

0.00kg

……………

When format setting is Ser.2, data is sent once when pressing M+, and print total value when pressing MC

S/N		WT		U/W		QTY

1		 64.3	0.53589	120		<-print single date when pressing M+

2		137.7	0.53592	257

3		168.8	0.53590	315

4		101.3	0.53591	189

5		++++			100		<-M+ manually enter the quantity without weighing

Sum				472.1g/981pcs	<-print total result when pressing MC

Auto: data is sent once when weighing result stables

When format setting is Ser.1, data is sent once when weighing result exceeds 10e and stabilized, but does not support print total

10.05kg

-10.05kg

6.00kg

0.00kg

……………

When format setting is Ser.2, data is sent once when weighing result stables, and print total value when pressing MC

S/N		WT		U/W		QTY

1		 64.3	0.53589	120		<-print single date when result stables

2		137.7	0.53592	257

3		168.8	0.53590	315

4		101.3	0.53591	189

5		++++			100		<-M+ manually enter the quantity without weighing

Sum				472.1g/981pcs	<-print total result when pressing MC

4.1.9 Turning On/Off Real Time Clock (OPTION)

Under this function, user can turn on or off real time clock function.

- Scale is off

- Press and hold TARE and press ON/OFF, display shows F1

- Continue pressing TARE until display shows F16

- Press ZERO to enter, and display shows rtC.oF

- Press MR to select between rtC.on or rtC.oF

- Press ZERO to confirm and exit the current function, display shows F17

- Press ZERO to enter the next function, or press M+ to exit. When M+ is pressed, display will show SAVE, user can either press ZERO for yes to save the changes made or press CL for not to save the changes.

When printing with RTC function enabled, it will look like:

When format setting is Ser.1:

2015/01/09	10:50:30	ST, NT,	100.01kg		<-net weight

2015/01/09	10:50:31	ST, NT,	100.01kg

2015/01/09	10:50:32	US, GS,  1000.01kg	<-gross weight

…………………………………..

When format setting is Ser.2:

2015/01/09 10:50:30

S/N		WT		U/W		QTY

1		 64.3	0.53589	120

2		137.7	0.53592	257

3		168.8	0.53590	315

4		101.3	0.53591	189

5		++++			100		<-M+ manually enter the quantity without weighing

Sum				472.1g/981pcs	<-print total result when pressing MC

4.1.10 Time and Date Setup

Use this function to setup year, month and date as well as time in 24 hours format.

- Scale is off

- Press and hold TARE and press ON/OFF, display shows F1

- Continue pressing TARE until display shows F17

- Press ZERO to enter, and display shows YEAr in the Unit Weight window with digits flashing in Total Count window

- Use the numeric keypad to enter the date information directly, and press M+ to confirm

- The setting currently engaging will be reminded in the Unit Weight window, repeat the above steps to complete setting on month, date, hour, minute and second

- When setting is completed, display shows F18 (or F20 if scale is in unlock condition)

- Press M+ to exit, display will show SAVE, user can either press ZERO for yes to save the changes made or press CL for not to save the changes.

- INSTRUCTION TO USE

The counting scale VGC/VOC is easy to use and user friendly. Everyone can start using the scale by following the below steps:

- The scale is placed on a solid surface

- Use adjustable feet and referring to the bubble level to make sure scale is in a level position

- Power on the scale and wait for the counting down

- Scale is in zero weight status and display showing zero indicator

- If zero indicator is not on, press ZERO to set the display to zero

- If zero weight status cannot be obtained, change the location and try again

  - WEIGHING

The standard procedure of a successful weighing is by placing the item on platter first and use the numeric keypad to enter the unit price. The gross weight of the item is displayed in the WEIGHT window and piece weight is displayed in the UNIT WEIGHT window.

When load is placed on platter and unit weight is entered, scale will show the total piece count of such load on TOTAL COUNT window.

To start the next weighing, remove the item from platter and put on the next item. Use the numeric keypad to enter the unit weight and total piece count will be shown in the TOTAL COUNT window.

  - HOW TO USE TARE FUNCTION

This scale equips with auto tare function (please refer to F9 of internal setting), auto-continuous tare, manual tare and preset tare.

- If auto tare function is enabled, weight will be automatically tared off the initial weight when place on the platter

- If auto continuous tare function is selected, weight will be automatically tared off once it stables until the full tare range

- If manual tare is needed, place the load/container on the platter first and press TARE key to tare off the weight. After weight is tared, NET indicator will appear indicating the weight showing on display is net weight

- If preset tare is needed, use numeric keypad to enter the tare value and then press TARE key. The display will show a negative preset tare value and both NET and PT indicators will appear

- The ZERO indicator points out the negative value shown means the platter is at zero weight status

- When the goods is placed on the platter with a positive value shown, this indicates the actual weight of the goods.

    - How To Clear Tare Function (auto tare, manual tare, preset tare)

- To clear auto tare, remove the initial weight/load from the platter and NET indicator will turn off. Press ZERO key to return to zero weight status if ZERO indicator does not appear.

- To clear manual tare, remove everything from platter and press TARE key, the NET indicator will turn off. Press ZERO key to return to zero weight status if ZERO indicator does not appear.

- To clear preset tare, press 0 on the numeric keypad first and then press TARE key to cancel preset tare. Please make sure scale is not engaged in any kind of tare before proceed to modify the pre-tare value.

NOTE:

It is regulated by OIML R76 that in dual range application, tare value will be cancelled automatically when returning from W2 to W1.

  - HOW TO USE THE COUNT FUNCTION

The result of total piece count is the result of calculation from weight being placed on the platter against with unit piece weight. There are two methods to obtain the unit weight:

    - Automatic Sampling

- Refer to the above steps about how to weigh and place a known quantity on platter

- Use numeric keypad to enter the known quantity and press QTY SET

- Scale will show the piece weight in UNIT WEIGHT window

    - Direct Entry

- Refer to the above steps about how to weigh and place a load on platter

- Use numeric keypad to enter the piece weight and press WT. SET

- The total quantity will be in TOTAL COUNT window

- To set another piece weight, remove the load and manually enter the unit piece weight and press WT. SET

- If a mistake a found when key in the value, press CL to start again. To cancel the unit weight, press 0 and then press WT.SET

  - HOW TO USE MEMORY FUNCTION

5.4.1 To Save To Memory

This scale comes with memory function and result can be transmitted through optional RS232 interface.

- Obtain the total count by either automatic sampling or direct entry

- Press M+

- UNIT WEIGHT window shows “Add. XX”, for “XX” is the number of transactions saved in memory, while the TOTAL COUNT window will show the accumulated result with TOTAL indicator on. Both information will remain on display for 1 second and return to normal weighing

- M+ indicator is switched on meaning there are data saved in memory

- Remove item from platter and repeat the above steps to save the next transaction

5.4.2 To Recall From Memory

- Press MR once to recall the total amount of transactions saved in memory, WEIGHT window shows the total accumulated weight, UNIT WEIGHT window shows “tot. XX”, for “XX” is the number of transactions saved in memory, while the TOTAL COUNT window will show the accumulated quantity with TOTAL indicator on

- Information on display will show for 3 second and return to normal weighing

5.4.3 To Delete From Memory

- Press MR to recall the accumulated data from memory

- Press CL after MR is pressed

- Memory is cleared and M+ indicator is off

- If any other key instead of CL is pressed after pressing MR, memory will not be deleted

  - HOW TO USE CHECK FUNCTION

The check function provides an audible warning towards the upper limit set by user for either total quantity or gross weight

- Press CHK and display will show CHECK on TOTAL COUNT window

- Use numeric keypad to enter either the upper limit for quantity or weight

- If the value entered is for total quantity, press QTY SET; if the value entered is for weight, press WT. SET

- If the value entered is incorrect, press CL to clear the value. If the intend function is wrongly pressed, restart the check function

- To cancel the check function, press CHK first and press digit key “0”. When the value 0 is flashing in the UNIT WEIGHT window, press either QTY SET or WT. SET to cancel the function which was previously engaged

  - AUTO PIECE WEIGHT ENHANCEMENT FUNCTION

This scale equips with auto piece weight enhancement function, which activates automatically when the piece weighing is obtained through 5.3.1 Automatic Sampling. This function provides a constant monitoring on the changes on the unit piece weight to obtain the most precise counting result. The updated unit piece weight will show on the UNIT WEIGHT window follows by a short beep.

In order to have auto piece weight enhancement function working properly, there are two conditions to meet:

- Unit weight is obtained by automatic sampling, the direct entry method will only have fix unit weight.

- After the initial sampling, the next load place on the platter must not exceeds 100% of the original count

  - HOW TO CONNECT RS-232 (D-sub 9 pin female connector)

- Refer to 4.1.8 to set the transmission protocol

- Data code: ASCII

Pin Assignment (PC end)

| Pin | Assignment |
| --- | --- |
| 1 | DCD |
| 2 | RXD |
| 3 | TXD |
| 4 | DTR |
| 5 | GND |
| 6 | DSR |
| 7 | RTS |
| 8 | CTS |
| 9 | NC |

Scale end

PC end

2

2

Only need to connect Pin2, 3 and 5

3

3

5

5

- DAILY CARE AND MAINTENANCE

-Do not attempt to open the scale or try to repair except authorized technicians only

-Always aware of the operating environment to avoid excessive moisture and extreme temperature

-Use original factory supplied switch adaptor for charging only

-Recharge the battery immediately when the battery level indicator starts flashing to avoid permanent damage to the rechargeable battery

-Do not put anything on the platter after weighing completes to avoid causing damage to the load cell

-If cleaning is needed, use damp cloth to clean the scale body. Stainless steel platter can be removed for cleaning purpose

-If cleaning agent is needed, please use mild detergent and avoid excessive water. Do not use any chemical or alcohol for cleaning purpose

-If you are uncertain about how to properly maintaining your scale to the optimal condition, please contact your dealer for further assistance

APPENDIX:

ERROR CODES

| ERROR CODE | DESCRIPTION | SOLUTION |
| --- | --- | --- |
| Err 0 | Flash memory accessing error | Change MCU, contact dealer for assistance |
| Err 1 | A/D accessing error | Check A/D, contact dealer for assistance |
| ------ | No load cell signal/load cell reading error | Check load cell and connection |
| Err 2 | Load cell wiring error | Check load cell wiring pin assignment |
| Err 3 | Offset value too low/too high |  |

DISPLAY SEGMENT DEFINATION